Key

CIED – Cardiac Implantable Electronic Devices

PPM – Pacemaker

CRT-P – Cardiac Resynchronisation Therapy with a Pacemaker

ICD- Implantable Cardiac Defibrillator

CRT-D – Cardiac Resynchronisation Therapy with a Defibrillator

Key

CIED – Cardiac Implantable Electronic Devices

PPM – Pacemaker

CRT-P – Cardiac Resynchronisation Therapy with a Pacemaker

ICD- Implantable Cardiac Defibrillator

CRT-D – Cardiac Resynchronisation Therapy with a Defibrillator

Cardiac Ring Magnet SOP Flowchart for Skin Surgery

Use of Magnets for Patients with Cardiac Implantable Electronic Devices (this guidance relates to the BHRS/BSDS Oct 23 guidelines for CIEDs).

Cardiac Ring Magnet SOP Flowchart for Skin Surgery

Use of Magnets for Patients with Cardiac Implantable Electronic Devices (this guidance relates to the BHRS/BSDS Oct 23 guidelines for CIEDs).

No

No

No magnet required

No magnet required

Yes

Yes

Does the patient have a CIED?

Does the patient have a CIED?

PPM/CRT-P

PPM/CRT-P

ICD/CRT-D

ICD/CRT-D

No magnet required

(follow steps in box 3 below for monitoring & cautery).

No magnet required

(follow steps in box 3 below for monitoring & cautery).

Surgical site above umbilicus

Surgical site above umbilicus

Surgical site below umbilicus

Surgical site below umbilicus

Magnet required

Magnet required

1

1

## Booking & Scheduling Checks

- Tick “Magnet Required" on booking form

- Surgical coordinator to highlight magnet need at start of list

- Confirm magnet availability before patient arrives in theatre

## Booking & Scheduling Checks

- Tick “Magnet Required" on booking form

- Surgical coordinator to highlight magnet need at start of list

- Confirm magnet availability before patient arrives in theatre

>

>

>

>

2

2

>

>

>

>

31

31

4

4

5

5

1

1

## Magnet Storage & Return

- Stored in locked store cupboard

- Return to same location post-use

- Clean if necessary

## Magnet Storage & Return

- Stored in locked store cupboard

- Return to same location post-use

- Clean if necessary

## Monitoring and Cautery Use During Procedure

- Use bipolar diathermy only at lowest setting possible with short bursts (<5 seconds, intervening pauses)

- Monitor patient during cautery through both verbal response and pulse oximetry

## Documentation

In operation notes and clinic letter include:

- Type of CIED

- When and where magnet was applied

- When it was removed

- Monitoring during the procedure

(e.g. "Patient alert, clinical and pulse oximetry monitoring performed")

## Documentation

In operation notes and clinic letter include:

- Type of CIED

- When and where magnet was applied

- When it was removed

- Monitoring during the procedure

(e.g. "Patient alert, clinical and pulse oximetry monitoring performed")

## Intraoperative Magnet application

When to apply:

- Just prior to the use of bipolar diathermy

How to apply:

- Place ring magnet inside a sterile glove

- Apply directly on top of CIED site

- Secure using two large Hypafix sheets

When to remove:

- After bipolar diathermy use is complete

## Intraoperative Magnet application

When to apply:

- Just prior to the use of bipolar diathermy

How to apply:

- Place ring magnet inside a sterile glove

- Apply directly on top of CIED site

- Secure using two large Hypafix sheets

When to remove:

- After bipolar diathermy use is complete

For any uncertainty or unresolved intraoperative issues, contact the on-call Cardiology team

For any uncertainty or unresolved intraoperative issues, contact the on-call Cardiology team